## Slide 1
~1Hz pulse rate (IPP=1 sec) in FreeRun Mode seems good for monitoring stimulus thresholds
PULSE SIGN
Biphasic = +
Baseline = Zero
DURATION
MULTIPLIER = 10-3 recommended so that the scroll wheel values are in milliseconds.
SCROLL WHEEL = 3 to 0.01 msec
RANGE SELECTOR = 10 Volts
(DO NOT USE Amps)
PULSE AMPLITUDE = 0-9 Volts
(DO NOT EXCEED (VOLTS!)
Start with 3 msec duration and 0 V.
Step 1: Slowly increase the amplitude until above spike threshold.
Step 2: Decrease duration until below spike threshold.
Repeat steps 1 and 2, recording the duration-amplitude pair at each spike threshold. Stop once you reach a short duration for which spike threshold would >9 V. DO NOT EXCEED 9V.
Stimulus output is a floating source (use ADC NRSE 10V range)

## Slide 2
I recommend using MANUAL trigger mode because of all the knob and wheel spinning that needs to happen to change IPI between trials
PULSE SIGN
Biphasic = +
Baseline = Zero
For all pulses, set AMPLITUDE above spike threshold at a Pulse DURATION of 200 microseconds (0.2 msec)
Stimulus output is a floating source (use ADC NRSE 10V range)
inter-pulse-interval on each trial is controlled by PERIOD. Experimentally range: 100 - 1 msec.
NUMBER OF PULSES per trial is controlled by the TRAIN BURST WIDTH. Set the burst width to between 1 to 2 times the IPI to get 2 pulses per trial. If burst width exceeds 2*IPI, then there will be more than two pulses per trial.

## Slide 3
You can either trigger each trial manually so that you test discrete sets of ipi instead or you can ‘repeat’ for a continuous range.
Must be in TWIN PULSES mode
If using REPEAT mode, I recommend ~1Hz IPI
DELAY is the time between pulses. Experimentally range: 100 - 1 msec
For all pulses, set AMPLITUDE above spike threshold at a Pulse DURATION of 200 microseconds (0.2 msec)
Stimulus output is a floating source (use ADC NRSE 10V range)